## Sheet1
| example values/format | min | max | variable description
CSTA1 | 1338000 | CC defined | CC defined | study identifer entered into the dicom header of the CT system
Subjid | J543376 | CC defined | CC defined | alpha / name code entered into the dicom header of the CT system
CSTA2 | 20010413 | 2002-12-31 | TBD | date of CT scan in dicom header
CSTA3 | smp |  | ldu,rch,smp | image analyst identifier for heart 1 (aka reader ID =ldu,rch,smp)
CSTA4 | 2 | 0.5 | 3 | size of region of interest used for calibration
CSTA5 | 3 | 0 | 3 | number of calibration standards, typically 3
CSTA6 | 0 | 0 | 0 | known value of rod 1 = 0 mg/cc
CSTA7 | 75 | 75 | 75 | known value of rod 2 = 75 mg/cc
CSTA8 | 150 | 150 | 150 | known value of rod 3 = 150 mg/cc
CSTA9 | -11 | -32 | 10 | measured value [CT numbers or H.U.] of rod 1
CSTA10 | 79 | 58 | 100 | measured value [CT numbers or H.U.] of rod 2
CSTA11 | 165 | 144 | 186 | measured value [CT numbers or H.U.] of rod 3
CSTA12 | 1.1112 |  |  | slope of regression fit
CSTA13 | 11.558 |  |  | intercept of regression
CSTA14 | 0.99876 |  |  | least squares
CSTA15 | 0 | 0 | 200 | no. of lesions left main
CSTA16 | 0 | 0 | 200 | no. of lesions left anterior descending
CSTA17 | 0 | 0 | 200 | no. of lesions left circumflex
CSTA18 | 0 | 0 | 200 | no. of lesions right coronary
CSTA19 | 0 | 0 | 200 | no. of lesions all coronary (does not include others)
CSTA20 | 0 | 0 | 5000 | Agatston score left main
CSTA21 | 0 | 0 | 5000 | Agatston score left anterior descending
CSTA22 | 0 | 0 | 5000 | Agatston score left circumflex
CSTA23 | 0 | 0 | 5000 | Agatston score right coronary
CSTA24 | 0 | 0 | 5000 | Agatston score all coronary (does not include others)
CSTA25 | 0 | 0 | 7000 | Volumeof lesions left main
CSTA26 | 0 | 0 | 7000 | Volumeof lesions left anterior descending
CSTA27 | 0 | 0 | 7000 | Volumeof lesions left circumflex
CSTA28 | 0 | 0 | 7000 | Volumeof lesions right coronary
CSTA29 | 0 | 0 | 7000 | Volume of lesions all coronary (does not include others)
CSTA30 | 0 | 0 | tbd | Mean of lesions left main (the max
CSTA31 | 0 | 0 | tbd | Mean of lesions left anterior descending
CSTA32 | 0 | 0 | tbd | Mean of lesions left circumflex
CSTA33 | 0 | 0 | tbd | Mean of lesions right coronary
CSTA34 | 0 | 0 | tbd | Mean of lesions all coronary (does not include others)
CSTA35 | 0 | 0 | 1500 | Calcium mass of lesions left main
CSTA36 | 0 | 0 | 1500 | Calcium mass of lesions left anterior descending
CSTA37 | 0 | 0 | 1500 | Calcium mass of lesions left circumflex
CSTA38 | 0 | 0 | 1500 | Calcium mass of lesions right coronary
CSTA39 | 0 | 0 | 1500 | Calcium mass of lesions all coronary (does not include others)
CSTA40 | 0 | 0 | 3 | Quality variable [ 0=NA, 1=poor, 2=fair, 3=good], analyst subjective rating of scan coverage, did scan include all of the coronary arteries
CSTA41 | 3 | 0 | 3 | Quality variable[ 0=NA, 1=poor, 2=fair, 3=good], analyst subjective rating of placement of QCT phantom
CSTA42 | 3 | 0 | 3 | Quality variable [ 0=NA, 1=poor, 2=fair, 3=good], analyst subjective rating of image noise level
CSTA43 | 2 | 0 | 3 | Quality variable [ 0=NA, 1=poor, 2=fair, 3=good], analyst subjective rating of streak artifacts related to motion unsharpness
CSTA44 | 3 | 0 | 3 | Quality variable [ 0=NA, 1=poor, 2=fair, 3=good], analyst subjective rating of mis-registration
CSTA45 | 3 | 0 | 3 | Quality variable [ 0=NA, 1=poor, 2=fair, 3=good], analyst subjective rating of other artifacts
CSTA46 | 4 | 0 | 5 | Quality variable [ 0,1,2,3,4,5 with 5 excellent], analyst subjective rating of overall series quality
CSTA47 | N | N | Y | heart selected for QC review yes/no
CSTA48 | N | N | Y | Flag variable for reader identified calcified plaque, not scoreable by minimum lesion critera
CSTA49 | N | N | Y | Flag variable for uninterpretable/studies which cannot be analyzed
CSTA50 | N | N | Y | Flag variable yes/no for reader identified ascending aorta calcified plaque
CSTA51 | 2 | 0.5 | 3 | size of region of interest used for calibration
CSTA52 | 3 | 0 | 3 | number of calibration standards, typically 3
CSTA53 | 0 | 0 | 0 | known value of rod 1 = 0 mg/cc
CSTA54 | 75 | 50 | 50 | known value of rod 2 = 75 mg/cc
CSTA55 | 150 | 100 | 100 | known value of rod 3 = 150 mg/cc
CSTA56 | -5 | -20 | 10 | measured value [CT numbers or H.U.] of rod 1
CSTA57 | 90 | 75 | 105 | measured value [CT numbers or H.U.] of rod 2
CSTA58 | 175 | 160 | 190 | measured value [CT numbers or H.U.] of rod 3
CSTA59 | 11.558 |  |  | slope of regression fit
CSTA60 | 0.99876 |  |  | intercept of regression
CSTA61 | 0 | 0 | 200 | least squares
CSTA62 | 0 | 0 | 200 | no. of lesions infrarenal abd. aorta
CSTA63 | 0 | 0 | 200 | no. of lesions left common iliac
CSTA64 | 0 | 0 | 200 | no. of lesions right common iliac
CSTA65 | 0 | 0 | 5000 | no. of lesions aorto-iliac
CSTA66 | 0 | 0 | 5000 | Agatston score infrarenal abd. aorta
CSTA67 | 0 | 0 | 5000 | Agatston score left common iliac
CSTA68 | 0 | 0 | 5000 | Agatston score right common iliac
CSTA69 | 0 | 0 | 7000 | Agatston score aorto-iliac
CSTA70 | 0 | 0 | 7000 | Volumeof lesions infrarenal abd. aorta
CSTA71 | 0 | 0 | 7000 | Volumeof lesions left common iliac
CSTA72 | 0 | 0 | 7000 | Volumeof lesions right common iliac
CSTA73 | 0 | 0 | tbd | Volume of lesions aorto-iliac
CSTA74 | 0 | 0 | tbd | Mean of lesions infrarenal abd. aorta
CSTA75 | 0 | 0 | tbd | Mean of lesions left common iliac
CSTA76 | 0 | 0 | tbd | Mean of lesions right common iliac
CSTA77 | 0 | 0 | 1500 | Mean of lesions aorto-iliac
CSTA78 | 0 | 0 | 1500 | Calcium mass of lesions infrarenal abd. aorta
CSTA79 | 0 | 0 | 1500 | Calcium mass of lesions left common iliac
CSTA80 | 0 | 0 | 1500 | Calcium mass of lesions right common iliac
CSTA81 | 3 | 0 | 3 | Calcium mass of lesions aorto-iliac
CSTA82 | 3 | 0 | 3 | Quality variable [ 0=NA, 1=poor, 2=fair, 3=good], analyst subjective rating of scan coverage, did scan include all of the aorto-iliacs
CSTA83 | 2 | 0 | 3 | Quality variable[ 0=NA, 1=poor, 2=fair, 3=good], analyst subjective rating of placement of QCT phantom
CSTA84 | 3 | 0 | 3 | Quality variable [ 0=NA, 1=poor, 2=fair, 3=good], analyst subjective rating of image noise level
CSTA85 | 3 | 0 | 3 | Quality variable [ 0=NA, 1=poor, 2=fair, 3=good], analyst subjective rating of streak artifacts related to motion unsharpness
CSTA86 | 4 | 0 | 5 | Quality variable [ 0=NA, 1=poor, 2=fair, 3=good], analyst subjective rating of mis-registration
CSTA87 | N | N | Y | Quality variable [ 0=NA, 1=poor, 2=fair, 3=good], analyst subjective rating of other artifacts
CSTA88 | N | N | Y | Quality variable [ 0,1,2,3,4,5 with 5 excellent], analyst subjective rating of overall series quality
CSTA89 | N | N | Y | aorta exam selected for QC review yes/no
CSTA90 |  |  |  | Flag variable for reader identified calcified plaque, not scoreable by minimum lesion critera
CSTA91 |  |  |  | Flag variable for uninterpretable/studies which cannot be analyzed
CSTA92 | Y |  |  | Evidence of medical/surgical intervention of CT images (yes/no)
CSTA93 | N |  |  | coronary artery bypass graft (y/n)
CSTA94 | N |  |  | pacemaker identified (y/n)
CSTA95 | N |  |  | coronary stent (y/n) -Note that the majority of stents may not be identified by simply reviewing the images and this is an underestimate
CSTA96 | N |  |  | cardiac valvular surgery (y/n)
CSTA97 | N |  |  | other surgery/intervention (y/n)
CSTA98 |  |  |  | Image slice number at which L4-5 is located in abdomen scan
CSTA99 |  |  |  | Image slice number at which abdominal fat measurements begin
CSTA100 |  |  |  | Image slice number at which abdominal fat measurements end
CSTA101 |  |  |  | Total fat volume for 50 mm of abdomen
CSTA102 |  |  |  | Visceral and intramuscle fat volume for 50 mm of abdomen
CSTA103 |  |  |  | Visceral fat volume for 50 mm of abdomen
CSTA104 |  |  |  | Total volume for 50 mm of abdomen
CSTA105 |  |  |  | Image slice number at which L4-5 is located in abdomen scan
CSTA106 |  |  |  | Image slice number at which abdominal fat measurements begin
CSTA107 |  |  |  | Image slice number at which abdominal fat measurements end
CSTA108 |  |  |  | Total fat volume for 10 mm of abdomen
CSTA109 |  |  |  | Visceral and intramuscle fat volume for 10 mm of abdomen
CSTA110 |  |  |  | Visceral fat volume for 10 mm of abdomen
CSTA111 |  |  |  | Total volume for 10 mm of abdomen
CSTA112 |  |  |  | Coronary scan phantom region of interest at LM 0 port mean
CSTA113 |  |  |  | Coronary scan phantom region of interest at LM 0 port standard deviation
CSTA114 |  |  |  | Abdomen scan phantom region of interest at bifurcation 0 port mean
CSTA115 |  |  |  | Abdomen scan phantom region of interest at bifurcation 0 port standard deviation